## Slide 1: In-core Instrumentation for MTR Experiments
Jean-François VILLARD
French alternative energies and atomic energy commission
Nuclear Energy Division – Reactor Studies Department
Cadarache – F-13108 St Paul Lez Durance, France
Joint ICTP/IAEA Workshop “Research Reactors for Development of Materials and Fuels for Innovative Nuclear Energy Systems”
6-10 November 2017, ICTP - Trieste, Italy

## Slide 2: In-core Instrumentation for MTR Experiments
Summary
General requirements
Conditions of measurement
Examples of in-core instrumentation : 3.1. Neutron and gamma flux
3.2. Temperature
3.3. Prospect : fiber optics
4. Review of MTR instrumentation around the World
PAGE 2

## Slide 3: 1. General requirements

## Slide 4: 1. General requirements
Illustration of measurement needs for material testing
PAGE 4
Objectives of irradiation tests :
Assess / verify / increase the precision of
behavior laws under irradiation
Integrated flux / dose
Studied property
(ex : fragility, cracking, corrosion, conductivity, etc.)
Influence parameter
(temperature, physicochemical conditions, etc.)
 All these parameters must be accurately measured !
Watch safety parameters
= check that some specified parameters stay in their acceptable range
(e.g. : temperature, pressure, etc.)
Monitor the experiments
damages
In-core measurements are essential to:

## Slide 5: 1. General requirements
PAGE 5
Main stakes for in-core instrumentation
Reduction of uncertainties in experimental conditions
- Neutrons / gamma flux
- Neutron spectrum
- Neutron dose
Better knowledge of nuclear heating
 design of experimental devices
Highly instrumented experiments
« Cook and look » 
Online measurement of experimental parameters
Material behavior
New materials, better representativeness
- Temperature
- Deformations (creep…)
Better characterization of thermo-hydraulic conditions
 temperature field / fluid flows
Fuel improvement and qualification
- Fission gas release
- Pellet-cladding interaction
- Accidental conditions (LOCA)
Radiation measurements
Physical parameter measurements

## Slide 6: 1. General requirements
Main reactor measurement needs and R&D programs
PAGE 6

## Slide 7: 1. General requirements
Yesterday:
Engineering approach
Mainly empirical characterization
Limited field of validation
Today:
Physician approach
Development and qualification of instrumentation modeling
Improvement in related nuclear data
Tomorrow:
Extended field of validation
Advanced signal analysis
Ways to develop and qualify in-core instrumentation are changing
PAGE 7

## Slide 8: 2. Conditions of measurement a. Physicochemical environment b. Effects of radiations c. Operational issues

## Slide 9
Physicochemical environment of in-core instrumentation
High temperature:
 LWR conditions : 300-400°C continuous (up to 1200-1800°C in transients)
 SFR conditions : 400-500°C continuous
 HTR-VHTR conditions, materials for fusion : 800-1200°C continuous
Various media:
 Pressurized water (+ bore and lithium)
 Liquid metals (NaK, Na…)
 High-temperature gas
Typical alloys used for in-core components
2. Conditions of measurement
a. Physicochemical environment
Type
Stainless steels
AISI 304 L
AISI 316 L
Nickel alloys (Inconel)
Alloy 600
Alloy 718
Zirconium alloys
Zircaloy 4
Zr-Nb
Others

## Slide 10
Main effects of nuclear radiations on sensors
 transmutations : composition changes
 damages :
- alteration of electric insulators
- wires breaking
- change in mechanical properties
 noise current (Compton and photoelectric effects)
 heating
 Precautions :
choice of materials
    form (metals, oxides, ceramics…)
    elements  nuclear properties
remove sensor from high-radiation areas when possible
in-situ calibration
2. Conditions of measurement
b. Effects of radiations

## Slide 11
Effects of transmutation on in-core sensors
Materials of standard thermo-elements | Si | Al | Cr | Mo | Ni | Pt | Mg | W | Re | Rh | Ir
(barn) neutron cross-section of the main isotope | 0,08 | 0,23 | 0,73 | 2,5 | 4,2 | 10 | 13,5 | 19,2 | 86 | 150 | 460
Illustration : effect on thermocouples
Change in wire composition
 change in thermoelectric response = drift of the signal
2. Conditions of measurement
b. Effects of radiations
the material X disappears under irradiation : N(t) = N0 e- t
Due to neutron capture

## Slide 12
Type of thermocouple | Drift in MTR due to irradiation
K (Ni-Cr / Ni-Al) | Not significant
C (W-Re5% / W-Re26%) | -1°C / 2.1019 n/cm2 (-0,9 °C / day)
S (Pt-Rh10% / Pt) | -1°C / 1,4.1019 n/cm2 (-1,2 °C / day)
Drift of thermocouple signal under MTR irradiation
W  Re + Os ( = 19 barn)
Re  Os ( = 86 barn)
Rh  Pd ( = 146 barn)
Important effect on high-temperature thermocouples (W/Re, Pt/Rd…)
2. Conditions of measurement
b. Effects of radiations
Effects of transmutation on in-core sensors
Illustration : effect on thermocouples

## Slide 13
Operational issues
High reliability (impossible maintenance on irradiated sensors)
High accuracy (higher scientific requirements. Example : in-core deformations must be measured with ~ µm accuracy)
Miniaturization (very narrow experimental devices : a few millimeters available for instrumentation)
Large distance between sensor and electronics (tens of meters)
2. Conditions of measurement
c. Operational issues

## Slide 14
Operational issues
Fragile instrumentation in difficult conditions of implementation
Repairing a mineral insulated cable
Damaged cables
2. Conditions of measurement
c. Operational issues

## Slide 15
Summary of in-core measurement conditions
High accuracy required to meet scientific and technical needs
High constraints related to reactor conditions:
  constraints due to irradiation (high neutron and gamma flux: material damages, composition changes, parasitic signal, etc.),
  constraints due to physicochemical conditions in experiments (high temperature, pressurized water, liquid metals, etc.),
  constraints due to integration (miniaturized sensors because of small size devices, long distance between sensors and electronics, etc.),
  constraints due to operation (high reliability requirements because of difficult or impossible maintenance or replacement of irradiated instrumentation).
PAGE 15
Innovation is necessary !
2. Conditions of measurement

## Slide 16: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux

## Slide 17: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Neutron and gamma measurements
Objectives :
  To measure and reduce the uncertainties on:
  Thermal and fast neutron flux
  Gamma flux and nuclear heating
State-of-the-art :
Neutron flux:
  Activation foil dosimeters and wires (post-irradiation analysis) + Unfolding techniques
  Fission chambers for thermal and fast neutron
  Self Powered Neutron Detectors
Gamma flux:
  Ionization chamber
  Self Powered Gamma Detector (with Bi emitter)
Nuclear heating:
  Calorimeter (Gamma Thermometer and Differential Calorimeter)
Developments :
  Simulation tools for FC and SPND signal
  Dosimetry for epithermal flux
PAGE 17

## Slide 18: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
= measurement of the radioactive activity of an irradiated material sample (dosimeter)
  Reference measurement for neutron flux and fluence
  High accuracy
  Selection of the neutron energy domain (thermal, epithermal, fast…)
  But post-irradiation analysis (no online measurement)
Activation dosimetry
Irradiation history
geometry model
Reaction rate, fluence…
Cross section
Activity
Irradiation
Transportation
Activity measurement (spectrometry  or X)
Dosimeters
Pic area
A(t) = Φ0 . N1 . σ1 . (1- e-λ.T) . e-λ.t

## Slide 19
Dosimeter
Dosimeter
Reaction
Reaction
Daughter
Daughter
Spectrum energy area covered (MeV)
Spectrum energy area covered (MeV)
days
days
days
years
hours
hours
hours
hours
hours
hours
a
3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Activation dosimetry
Short-life materials for dosimeters

## Slide 20
3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Activation dosimetry
Long-life materials for dosimeters

## Slide 21
Dosimeter location
3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
PAGE 21
Activation dosimetry
Illustration of dosimeter implementation : neutron fluence monitoring for material testing experiments

## Slide 22
3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Activation dosimetry
Illustration of dosimeter implementation : flux mapping
PAGE 22

## Slide 23: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
    Performs reactor dosimetry measurements for research reactors and for the Surveillance Program of French PWRs
     Measurement of  and X activity of solid samples
     Range : 0.1 Bq to 108 Bq / dosimeter
     Uncertainties : 1% à 5% (k=1)
PAGE 23
Illustration of measurement facility : MADERE Platform at CEA Cadarache
Activation dosimetry

## Slide 24: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Measurement of the current generated by fission reactions in a fissile material deposited on an electrode
An effective detector for measuring neutron flux
High linearity signal (10 decades)
Real-time online measurement
A modular detector
Choice of fissile isotope and mass deposit
Adaptation of technological parameters:
geometry, nature and pressure of gas...
A multi-information signal
Three operating modes depending on flux range
Neutron / gamma effective discrimination
PAGE 24
Fission chambers

## Slide 25: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
PAGE 25
Count rate
Pulse mode
 suppression
threshold
Campbell mode (MSV)
V² = KNQ²
V² : variance
Q : mean charge per event (Q < Qn)
N : mean event rate
  rejection
Current mode
> 5.1010 n.cm-2.s-1
Campbell mode reduces  contribution by a factor of 50% / 0.6% = 83 for 242Pu FC
and 10.5% / 0.4% = 26 for 235U FC
Fission chambers

## Slide 26: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
   Manufacturing specific fission chambers with various possible fissile deposits (mass/composition) : U, Pu, Th, Np, Am, Cm… and specific designs
   Imagery and controls (RX)
PAGE 26
Fission chamber manufacturing workshop at CEA Cadarache
Fission chambers

## Slide 27: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Application of fission chambers :
Fast Neutron Flux Measurement
Objective : Online measurement of the fast neutron flux (E>1 MeV) in typical Material Testing Reactor conditions : high neutron flux (1015 n.cm-2.s-1) generally dominated by thermal neutrons + high gamma flux
Principle : use of a specific fission chamber (242Pu) and specific acquisition and software
Range of interest
Reaction of interest 242Pu fission
Parasitic reactions
PAGE 27

## Slide 28: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Application of fission chambers :
Fast Neutron Flux Measurement
PAGE 28
2 patents

## Slide 29: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Application of fission chambers :
Fast Neutron Flux Measurement
FNDS is the first and unique acquisition system able to provide an online measurement of the fast neutron flux (E>1MeV) in typical MTR conditions
 FNDS proved its ability to measure online both thermal and fast neutron flux (consistency with activation dosimetry is better than 5%).
FNDS has been validated in 2015 after successive tests at the Belgium BR2 reactor and at the French ISIS reactor (FNDS signal was compared to reference thermal and fast neutron flux measurements using activation dosimeters)
PAGE 29

## Slide 30
Application of fission chambers :
Neutron profile measurements between fuel plates in OPAL reactor (2013)
Experimental rig
  3 x Ø1.5mm Fission Chambers
  Dedicated Experimental Rig:
  2mm thick aluminium plate attached to a 16 meter long pole
      Correct and repeatable positioning
      Precise manoeuvring within the core
  OPAL Reactor operated at stabilized power 100 kW
Ø1,5 mm Fission Chambers
for three simultaneous neutron profiles
Insertion between fuel plates
3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
PAGE 30

## Slide 31
Application of fission chambers :
Neutron profile measurements between fuel plates in OPAL reactor (2013)
  400 measurement points
  Thorough investigation of 5 / 16 fuel assemblies:
      One complete quadrant
      One fresh fuel assembly in another quadrant
  Good agreement with expected neutron flux shapes
Measurement Locations
3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
PAGE 31

## Slide 32: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Miniaturized Ionization Chamber
 online and real-time evaluation of gamma flux over a wide range of flux level.
 monitoring reactor power / following reactor SCRAMs (reactor shutdown with rapid insertions of control rods).
PAGE 32
        Pure gamma selectivity
        Large flux range and good linearity
        (up to a few 1014 γ.cm−2.s−1)
        Small dimensions (3mm diameter)
        No burn up effect

## Slide 33: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Self-powered neutron detectors (SPND)
+
=
PAGE 33
  On line neutron or gamma flux measurement/monitoring
  Robust design and small dimensions
  Relatively easy implementation
  Acceptable cost
Basic Principle :
  Collection of electrons (or b-) created from nuclear interactions with materials
  (= self-generated current)

## Slide 34: 3. Examples of in-core instrumentation 3.1. Neutron and gamma flux
Main Characteristics :
  Usual emitter materials: Ag, Rh, Co, V, Pt, Hf… Gd, Bi (gamma)
  Typical direct current for 2.1014 n.cm-2.s-1: Ag ~ 0,8 mA , Rh ~ 1,6 mA
  Saturation time: from 30 ms to 30 min
  Depletion for 2.1014 n.cm-2.s-1: Ag ~ 3%/20 days, Rh ~ 5%/20 days
  Not a selective sensor : thermal and epithermal neutron sensitivity (+ gamma)
  Usual geometries:
Sensitive length from 1 to 10 cm
Overall diameter from 1,5 to 4 mm
Typical cable length from 10 to 20 m
Usual insulator material: MgO ou Al2O3
Precautions :
  Emitter depletion to be calculated at all time for (n,b-) detectors (irradiation history is necessary)
  Possible weak signal  measurement precautions ( < 1 nA for low reactor power)
  Correction factors (cable, gamma contribution,...) to be mastered
  Output signal is integral; processing not straightforward
   A multistep Monte Carlo calculation model has been developed and qualified for SPND signal analysis
Self-powered neutron detectors (SPND)
PAGE 34

## Slide 35: 3. Examples of in-core instrumentation 3.2. Temperature

## Slide 36
Temperature measurements
Objectives :
  Online measurement of materials or fuels temperature
  Typical range : 200-400°C (PWR materials) up to 1200°C (fuel)
  1000-1600°C for transient / study of incidental conditions
State-of-the-art :
  MIMS thermocouples : type K, N, C
  Expansion thermometers (LVDT)
  Melt wires, paint spots and SiC monitors (post-irradiation analysis)
Developments :
  High-temperature doped Mo/Nb alloy thermocouples
  In-situ calibration of thermocouples :
    Noise thermometry
    High-temperature fixed-point -cell
  Ultrasonic sensors
  Infrared pyrometer
  Distributed temperature measurements with Optical Fibers
PAGE 36
3. Examples of in-core instrumentation 3.2. Temperature

## Slide 37
Principle : Seebeck effect
An electromotive force is generated in an open circuit made of 2 different conducting materials, when their junctions are at different temperatures
Characteristics of in-core thermocouples :
Couples :
type K or type N (up to 1100 °C depending on wire diameter)
type C (fuel centerline temperature, for short irradiation time)
Sheath : Stainless steel (304L, 316L, 347)
Insulator : Al2O3, MgO, HfO2
Standard sheath diam. : 1 mm
Standard wire diam. : ~ 0,2 mm
Hot junction insulated from sheath
Length : tens of meters
3. Examples of in-core instrumentation 3.2. Temperature
Thermoelectric wires
Mineral insulation
Metallic sheath
PAGE 37
Thermocouples

## Slide 38
Thermocouples
Examples of implantation in MTR experiments
ERABLE sample holder
Tensile test specimens
CT specimens
Thermocouples
3. Examples of in-core instrumentation 3.2. Temperature

## Slide 39
Thermocouples
Examples of implantation in MTR experiments
CARDINAL sample holder
Location of specimens
for creep or growth study
Thermocouples
3. Examples of in-core instrumentation 3.2. Temperature

## Slide 40
Drilled pellets
3. Examples of in-core instrumentation 3.2. Temperature
Thermocouples
Examples of implantation in MTR experiments
for fuel centreline temperature measurements

## Slide 41
Expansion thermometer
Irradiation device structure
Fuel rod head plug
Plenum spring
Fuel rod
Tungsten rod (dia. 1,5mm)
Fuel rod bottom plug
Core support
Magnetic core
Elongation sensor (LVDT)
Measures the expansion of a metallic rod
+ robust
+ suitable for very high temperatures
integrated temperature
very intrusive
3. Examples of in-core instrumentation 3.2. Temperature

## Slide 42
Ultrasonic sensors
Measures the acoustic wave velocity in a metallic rod or wire
Notches  wave reflection (echoes)
    Detection of acoustic wave velocity changes between notches
 High-temperature measurements (> 2000°C)
 Distributed measurements
3. Examples of in-core instrumentation 3.2. Temperature

## Slide 43
Melt wires
Based on the detection of wire melting
 Easy implementation (no cable)
But - post-irradiation measurement
- only indicates that melting temperature has been reached
3. Examples of in-core instrumentation 3.2. Temperature

## Slide 44: 3. Examples of in-core instrumentation 3.3. Prospect : fiber optics

## Slide 45
Recent results : new fibers can survive in-core
Example : COSI experiment (OSIRIS reactor, France – 2006)
Location of the tested fibers
emission
reception
92 days of reactor operation
Dose > 1E20 nfast/cm-2
> 16 GGy
Irradiation of 12 recently developed optical fibers
3. Examples of in-core instrumentation 3.3. Fiber optics

## Slide 46
Favorable spectral region in the 800-1200 nm range
RIA measured losses < 10 dB  suitable multimode and single mode fibers exist for in-core applications
Better choose measurement systems that do not rely on light intensity
 interferometry (Fabry-Perot, Bragg gratings…)
3. Examples of in-core instrumentation 3.3. Fiber optics
Recent results : new fibers can survive in-core
Some results from COSI experiment

## Slide 47
Example of fiber-based in-core sensor : optical extensometer
Nuclear-grade elongation Fabry-Perrot type sensors
Adaptation to high T (400°C), with HT resistant fiber (Cuball, Al, Au coating)
Better sensor-sample sealing
Increased accuracy (1µm/100µm) and temperature compensation
PAGE 47
3. Examples of in-core instrumentation 3.3. Fiber optics
 Low drift at high temperature and high flux

## Slide 48
PAGE 48
In-core optical measurements
Fabry-Perot sensor is now in a phase of industrial transfer
Other developments : Bragg Grating sensors
= distributed measurements of temperature / deformation
Unstressed fiber
stressed fiber
Bragg gratings
Spectral shift
Wavelengh
On-going tests at the MITR
Scheduled tests at the ATR and BR2
3. Examples of in-core instrumentation 3.3. Fiber optics

## Slide 49: 4. Review of MTR instrumentation around the World

## Slide 50
4. Review of MTR instrumentation around the World

## Slide 51
4. Review of MTR instrumentation around the World

## Slide 52
4. Review of MTR instrumentation around the World

## Slide 53: Nuclear Energy Division Reactor Studies Department Experimental Physics Section Instrumentation Sensors and Dosimetry Laboratory
French alternative energies and atomic energy commission
Cadarache | F-13108 Saint-Paul-lez-Durance | France
T. +33 (0)4 42 25 79 62 | F. +33 (0)4 42 25 78 76
Etablissement public à caractère industriel et commercial
RCS Paris B 775 685 019